# Checklist for Safe Use of Biological Safety Cabinets

Office of Research Compliance, Integrity and Safety

This checklist is a template you can edit and modify as necessary to incorporate your laboratory-specific standard operating procedures (SOPs). Use this checklist as a daily reminder of the activities/tasks needed for safely working in a biological safety cabinet (BSC), as a training tool, or for operational protocols.

# Biological Safety Cabinet (BSC) Checklist

## Preparing for Work in a BSC

- Put on PPE according to your laboratory SOP

- Turn UV light off (if used).

- Turn fluorescent light on.

- Turn the blower motor on, allow it to run for 4 minutes (or Manufacturer’s recommended time) to purge the BSC of particulates. Some cabinets may alarm until purge is complete.

- Verify proper sash height and that sash alarm is on.

- Verify drain valve underneath the cabinet is closed (valve handle is perpendicular to valve body).

- Check cabinet’s certification sticker expiration date is within 1 year.

- Record the pressure differential gauge reading (if present), compare it against the calibration set point.

- Verify inward airflow according to your laboratory SOP (e.g., tissue test).

- Schedule uninterrupted work time, if possible.

- Decontaminate all surfaces of the cabinet, according to your laboratory SOP

- Collect all materials needed for the work.

## Safe Use of a BSC

- Place absorbent plastic-backed material to protect the work surface, if required by your lab SOP

- Wipe the external surfaces of any equipment or supplies that you will need to place in the BSC

- Place materials as far back in the cabinet as practical without blocking front and rear grilles.

- Separate clean/sterile items from dirty/potential contaminated items inside the BSC

- Work from clean to dirty to minimize cross-contamination.

  - One person at a time should work in a cabinet. If two people need to work in a B.S.C, use a 6’ cabinet and document the protocol-driven need in a risk assessment completed by the Lab Supervisor prior to starting work.

- Move your arms slowly in and out of the BSC Do not move arms in sweeping, sideways movements.

- If using the vacuum system, protect the vacuum line with a filter.

- Adjust the stool/bench height so your face is above the bottom of the sash and your arms enter the cabinet with elbows at 90° angles, armpits level with the bottom of the sash. Use a footrest if your feet do not touch the floor.

- Discard all waste in a biohazard bag inside of the B.S.C.

## After Completing Work in a BSC

- Do not turn blower motor off while removing items and decontaminating the cabinet.

- Close and surface decontaminate all containers and lab materials before removal.

- Surface decontaminate the exterior of the biohazard waste bag/container before removal and disposal.

- Decontaminate cabinet interior, including sidewalls, back wall, inside of sash and work surface.

- Remove the gloves you were using in the BSC and dispose of them according to your laboratory SOP

- Turn fluorescent light and blower motor switches off.

- If required by your lab SOP, turn UV light on. Keep the sash closed while the UV light is on.

- Wash your hands.

This job aid is a component of the free, on-demand CDC training course “Fundamentals of Working Safely in a Biological Safety Cabinet.”

v.23078